## Slide 1: Microscope Systems with Lasers
Confocal Microscopes
Usually sold as Class 3B systems because beam is accessible at the stage (If in doubt the user manual for the system should declare the laser product classification of the set up as sold).
Exposure is unlikely during normal use; So most (but not all) confocals have been granted some exemptions from Class 3B requirements but the following is still required:
- Users must have taken the Laser Safety Awareness Course (THS42e).
- Users must acquire training from an expert on the specific microscope system.
- Laser Risk Assessment and Local Rules (SOPs) documents must be completed for the specific system and must be read by all users.
- A list of authorised users must be kept.

## Slide 2: Microscope Systems with Lasers
Confocal Microscopes
The laser power must be key operated and the key should be removed and stored safely when the lasers are switched off (where this is not practical; e.g. in a core facility; then provide password
access to the software laser controls instead).
All confocals should be configured so that laser light can not enter the eyepiece (either by the design of the optical path or by using filters or interlocks).
Inverted confocals should either have fixed condenser arms or have interlocks that prevent the laser entering the objective when the condenser arm is tilted backwards.

## Slide 3: Microscope Systems with Lasers
Confocal Microscopes
The lowest laser power practicable should be used for excitation (to prevent sample damage / bleaching as well as reducing the danger to the user!)
Environmental control chambers will normally enforce a safe working distance from the beam at the sample
- Take extra care if one isn’t in place.
- Do not remove it if present!
- Clear chambers will not protect from specular reflections!
Specular reflections from the beam at the stage are hazardous!
To avoid these:
- Do not adjust the sample or change lenses while the laser is scanning.
- Do not place metalized or reflective objects in the sample area while scanning.
Do not remove panels / optical fibres as they are not necessarily interlocked and the laser power could be greater here than at the sample stage.
Diffuse scatter of laser light by the sample is not usually dangerous but avoid staring at the sample while it is scanning as a precaution.

## Slide 4: Microscope Systems with Lasers
Multiphoton / STED microscopes
MP and STED lasers are often Class 4 even at the sample stage.
In addition to the requirements for Class 3B microscopes:
- Everybody with access to the lab must take the Laser Safety Awareness Course (THS42e).
- Users should take the Amsler Grid self test and refer themselves for laser worker medical screening if necessary.
- A laser worker checklist should be completed.
Even diffuse scatter from the sample can cause serious, irreversible damage to the eyes and skin.
-Take extra care not to stare at the sample while the laser is scanning.
If possible a Black-out chamber should be fitted around the stage on Class 4 microscope systems.
If a black-out chamber is not present then
  the microscope should be covered with a black out cover during scanning.
  If the system can not be covered then all lab users should wear appropriate laser goggles.
  If a clear chamber is fitted on an inverted microscope then a shield should be fitted to the condenser to prevent specular reflections (and most of the diffuse scatter) from reaching the user.

## Slide 5: Microscope Systems with Lasers
Beam Alignment and System Reconfiguration
During laser replacement, system upgrades / reconfiguration and laser re-alignment it might be necessary for interlocks to be overridden and for enclosures to be opened up.
If this is the case then
  The lowest laser power should be set for alignment
  If lasers are being aligned down optical fibres then the open end should be clamped to the bench facing away from lab users.
  Specific SOPs for this task should be followed.
  Signs warning of the open beam work / alignment should be placed on a barrier across the lab door
  Appropriate laser goggles should be worn by everybody in the lab (if required by the SOPs).
University personnel involved in beam alignment and maintenance (and anybody involved in writing a laser risk assessment) must take an Advanced Laser Safety Training Course in addition to the THS42e Awareness Course.
If a technician is brought in to do the work then they will follow their own SOPs but warning signage should still be put in place across the lab entrance.
  - The technician must complete a Laser Permit to Work form and they must sign off to confirm that all enclosures have been replaced and interlocks have been reactivated before they leave.

## Slide 6: Microscope Systems with LEDs
LED Illumination
Many microscope systems now use high powered LEDs for widefield fluorescence illumination.
LEDs were taken out of the laser classification system about ten years ago and are now classified in a similar way to bright lamps.
Unlike with lasers there is no legal requirement for companies to label LED products with a safety classification.
- It is therefore best to measure the accessible power output of LED based systems to make sure that the
levels are safe.
Most LED systems attached to microscopes (e.g. CoolLED / Lumencore) are RG2 but some of the more powerful UV and IR lamps might be considered RG3.

## Slide 7: Microscope Systems with LEDs
LED Illumination
Reflections of the LED light at the sample stage are unlikely to be dangerous but it is best to avoid staring at the scattered LED light at the stage as a precaution.
LED power should be set as low as practicable for the experiment.
Direct viewing of the LED excitation light should be avoided and fibre optics and enclosure panels should not be removed while the LEDs are switched on.